## Slide 1: High Gradient RF Lab at IFIC
Daniel Esperante Pereira
On behalf of the Group of Accelerator Physics (IFIC)
C. Blanch, M. Boronat, D. Esperante, J. Fuster, B. Gimeno, D. González, A. Vnuchenko

## Slide 2: Outline
Overview and motivation
The IFIC-HGRF lab, design aspects and status:
  The XBOX inspiration
  The HPRF subsystem
  The LLRF subsystem
  New developments
Summary & plans

## Slide 3: The IFIC-HGRF Lab: Overview
Aim:
  High-gradient normal conducting RF cavities research topics at S-Band (2.9985 GHz)
RF Frequency: 2.9985 GHz
Repetition rate: 400 Hz
Up to now ~20 MV/m => With high-gradient technology up to 100 MV/m
Pulse length: 5 µs
Pk power:
15 MW

## Slide 4: The IFIC-HGRF Lab: Overview
Aim:
  High-gradient research topics at S-Band (2.9985 GHz)
Technical specifications:
  S-Band (2.9985 GHz)
  Pulsed High Power RF (HPRF):
    2 x pulsed power klystron: to 7.5 MW, 5 us pulse, 400 Hz
    High power waveguide RF network that allows power combining: enables to test 2 structures at a time at up to 15 MW, 5 us pulse, 200 Hz repetition rate
  Low level RF (LLRF): real-time control system with fast system interlock based on Ni-PXI acquisition system
  Running on Ultra-High vacuum (10-9 mbar)

## Slide 5: Motivation
Funding for research high-gradient/high performance RF topics in S-Band:
  Granted funding by the Spanish government coming from ERDF funds
  Got funds for a collaboration project between IFIC and CERN-CLIC => KE project, apply HG to medical physics
  Recently:
    Marie Curie – IF: D. Esperante, grant to develop the system and upgrade it
    Marie Curie – ITN-OMA: A. Vnuchenko. Optimization of Medical Accelerators Innovative Training Network (H2020)
Build a facility to open the possibility of several High-Gradient RF studies in S-Band:
  High-Power performance: of different prototype designs to determine the right direction towards higher gradients at acceptable breakdown rates
  Breakdown rate dependencies: with gradient and pulse length, to test the experimental power laws and develop physical models
  Breakdown localization: techniques to monitor the performance of the structure and better understand the BD formation process and precursors
  Surface field emission: current measurements
  Development of alternative diagnostics and analysis techniques: to follow further breakdown studies and high-power phenomena
  Design and development of RF components and technology (mostly in S-Band)

## Slide 6: System level simplified scheme
System level simplified scheme
Simplified, NOT SIMPLE!! 
WP3
WP2
WP4
Output data for
character-ization and analysis

## Slide 7: Outline
Overview and motivation
The IFIC-HGRF lab, design aspects and status:
  The XBOX inspiration
  The HPRF subsystem
  The LLRF subsystem
  New developments
Summary & plans

## Slide 8: The HPRF subsystem: XBOX inspiration
Klystron + modulator
Directional coupler
3 dB hybrid combiner
RF load
7.5 MW x 5µs x 400 Hz
RF/vacuum pumping port
DUT
with Faraday cups and Pumping ports
BUNKER
15 MW x 5µs x 200 Hz
Preamp
The XBOX’s are X-band (11.994 GHz). The Val-HGRF lab is S-Band (2.9985 GHz)
Try to keep as much as possible the same of hardware and software. (This is the most cost-effective solution: optimize the development time and manpower )
Pulse compressor

## Slide 9: Outline
Overview and motivation
The IFIC-HGRF lab, design aspects and status:
  The XBOX inspiration
  The HPRF subsystem
  The LLRF subsystem
  New developments
Summary & plans

## Slide 10: The HPRF subsystem
7.5 MW x 5µs x 400 Hz
BUNKER
15 MW x 5µs x 200 Hz
SSPA
Modulator+Klystron
Waveguide components
Status summary:
  Modulator: Jema. 1st modulator test with Klystron done by mid-September
  Klystron: CPI. Factory tests OK. Integration test with modulator done
  Waveguide components: MegaInd. Vacuum and RF tested
  Solid-state power amplifiers: MicroWave Amplifiers. Tested

## Slide 11: The HPRF subsystem: Amplifiers
Klystron VKS8262G1:
  Provider: CPI
  LIL output port flange for UHV
  Peak power: 7.6 MW @145 kV & 105 A
  Tp = 5 us, 400 Hz
  Status: RF tests in CPI. Now in JEMA facility (integrated into the modulator)
SSPA (Solid-state Power Amplifier):
  To drive the Klystron
  Solid-state: MicrowaveAmplifiers:
    AM10 S-Band 400W Power Amplifier.
  Acceptance tests done
  Status: tested at IFIMED, and one to be used at CERN soon

## Slide 12: HPRF: Modulator (1)
This is the one of the key elements of the system.
Solid-state switching modulator. Specs:
  Provider: JEMA (Lasarte-Oria)
  Vpeak, typical: 145 kV
  Ipeak, typical: 105 A
  Tp = 5 us, 400 Hz
  Voltage flatness: <=+/-0.25%
  Stability, pulse to pulse: <= 0.1%
  Risetime (10->90%): <=2us
  Falltime (90->10%): <=2us
  Adjustable voltage: (45kV - 150kV)
  Status: Integration with Klystron done (see next slides)

## Slide 13: HPRF: Modulator (2), Visit on September
Visit to Jema on the 21st of September: G. McMonagle (CERN), O. Sablic (CPI) and my self (IFIC).
  Purpose: integration of modulator with Klystron
  Outcome very successful:
    The tests were OK, despite some shortcomings:
      All interlocks OK
      Klystron pulsed at 100 Hz (No RF)
      Pulse to pulse stability OK
      Tests at 400 Hz with load OK
      Some few fixes to be done, already implemented
      The flat top of the pulses has some ripple (to be understood)
      Pulse width higher than expected. Rise time is higher than 2 us, while fall time looks good
Next steps:
  Modulator shipped to IFIC, integrate it and start pulsing: 13 March 2018

## Slide 14: The HPRF subsystem: waveguide components
Klystron output
Directional coupler
3 dB hybrid combiner
RF load
RF/vacuum pumping port
DUT
with Faraday cups and Pumping ports
WR284 waveguides

## Slide 15: The HPRF subsystem: lab integration
Accelerating structures
Devices under test inside the bunker
Modulator+ 7.5 MW Klystron units
3dB hybrid
Bunker
Accelerating structures
The bunker is in place.
The supporting structure as well.
New integration plan to be done due to modulator size mismatch.

## Slide 16: The HPRF subsystem: NEW lab integration
New integration plan designed due to modulator size mismatch
Less free space than before but things still fit

## Slide 17: HPRF installation: Modulator and waveguides
The bunker:
  The bunker (1), the crane (2) and the supporting structure (3)
  Waveguides (4)
  Rack for vacuum components (5)
  Modulator and Klystron connected (6). Some lessons learnt and fixes applied.
1
2
3
4
5
6
6

## Slide 18: Outline
Overview and motivation
The IFIC-HGRF lab, design aspects and status:
  The XBOX inspiration
  The HPRF subsystem
  The LLRF subsystem
  New developments
Summary & plans

## Slide 19: The LLRF subsystem
7.5 MW x 5µs x 400 Hz
- Slow Amplitude only. Log det. @250MSPS  Reflected signals/machine protection (x6)
- Phase & amplitude @62.5 MHz @ 250MSPS  Phase but not speed is key
- Phase & amplitude @62.5 MHz. @250MSPS  Structure RF Signals
-Faraday Cup Signals (x4)
BUNKER
15 MW x 5µs x 200 Hz
BDR studies:
20 ns time resolution needed
RF diagnose signals:

## Slide 20: LLRF subsystem components: status (1)
LLRF
Calibration power meter
Log detector
Down-mixer
Environmental monitoring
High Power RF:
Solid state amplifier
Klistron+modulator
RF network
Device under test
National Instruments
PXI real-time system:
RF generation
RF signals acquisition
Signal processing
External systems control, environmental monitoring and interlock
Ext. systems:
Cooling
Vacuum
Access
Modulator interlocks
Control, monitoring and interlocks
RF generation
RF diagnosis
Environmental signals
Intfce Intlck Box
UNDER CONSTRUCTION!!

## Slide 21: System layout: subsystem interconnect level
Explanation for kids:
Imagine this is smurfland

## Slide 22: A name for the lab
We should coin a name to de Lab.
It’s difficult to compete with a name so notorious as the XBOX…
So here my proposal…
S-Band Multi-Use Radio-Frequency Lab
SMURF Lab

## Slide 23: The LLRF system: status (2)
Down-mixer: Fully assembled and tested!!
16 channel downmixing system to move from 3 GHz to 62.5 MHz and get 4 samples per RF period and perform a “quick” IQ demodulation
One built and being used at CERN for the testing of the KT structure
One replica built at IFIC
PLL synthesizer
Power splitter
Mixing channels
PLL synthesizer
Power splitters
Mixing channels
Notes: Already explained to slides before the purpose of the different devices.

## Slide 24: The LLRF system: status (3)
Log amp: Under development, but not critical for startup!!
Needed to detect the the power of reflected signals and fast interlock if the power is too high in order to for Klystron protection:
8 channels, 1-10 GHz detection bandwidth
45 dB dynamic range and 20 ns rise/fall time
It consists of:
8 Log-detector PCB (inside a shielded box)
An 8 channel amplifier and signal conditioning PCB Outputs analog power levels and digital interlocks
Electrical tests done
Mechanical components built
Assembly process ongoing
Single channel log-detector
Notes: Already explained to slides before the purpose of the different devices.

## Slide 25: The LLRF system: status (5)
Interlock system:
Electrical engineering Bachelor degree thesis, A. Delgado Llop.

## Slide 26: The LLRF system: status (4)
Trigger-interlock box:
It adapts the input and output voltage signal levels from the NI-PXI to the rest of the setup
Fully assembled and tested!!
RF distributor:
It splits the RF signals coming to the different stages: down-mixers, log-detector, power calibrations
Built and fully characterized with the VNA by the Bsc. student
Other elements:
Cabling installation and calibration missing. Need to wait for modulator installation
Full integration to be done

## Slide 27: The LLRF system: status (5)
PXI and LabVIEW:
Chasis | 1 |  | NI PXIe-1085, 18 slots
Controller PXI | 1 |  | NI PXIe-8135
Trigger+interlock card | 1 |  | NI 6583 + FPGA
RF Generator (4.4 GHz) | 2 |  | NI 5793 + FPGA
4 ch 250 MSPS ADC | 6 |  | NI 5171. 4xAC(RF signals) + 1xDC (Faraday cup)
RS232, RS485 gauge det. | 2 |  | NI PXI-8432/8 (RS232) and NI PXI-8433/4 (RS485). Vacuum gauges
Slow monitoring | 1 |  | NI 6363. Voltage signals like vacuum, temperature and flowmeter
Fast acq 4 ch 2.5 GSPS | 1 |  | NI 5160. Not needed for S-band: cell bining resolution of 20 ns

## Slide 28: Outline
Overview and motivation
The IFIC-HGRF lab, design aspects and status:
  The XBOX inspiration
  The HPRF subsystem
  The LLRF subsystem
  New developments
Summary & plans

## Slide 29: New RF developments
Spherical compact RF Pulse Compressor. Joint development:
  C. Serpico and S. Nuaman from Elettra are doing all the RF design with support from A. Grudiev (CERN)
  IFIC will take care of the vacuum simulations and mechanical design
140 mm
347 mm
200 mm
595 mm

## Slide 30: Outline
Overview and motivation
The IFIC-HGRF lab, design aspects and status:
  The XBOX inspiration
  The HPRF subsystem
  The LLRF subsystem
  New developments
Summary & plans

## Slide 31: Summary
  RF Lab construction:
  Lab building construction is ready
  Lab design concept performed
  Commercial equipment reception and acceptance mostly done
  Redesign of some of the LLRF elements done
  Vacuum system tests done
  Modulator commissioning with Klystron done
  To do:
    Few components still to be finalized
    Control software being readapted
    Cooling system and demineralizer installation under final commissioning
    Entering in commissioning phase, first modulator pulsing by the end of the year
    Take care of the pulse compressor
  The new IFIC RF labs is on benefit of an open collaboration between IFIC, CERN, others for High-Gradient developments
  We’ll soon have and RF lab with high-power / high-gradient / high-rate testing capability

## Slide 32: Acknowledgements
Former and present team group members and students: A. Faus-Golfe, J. Giner, C. Blanch, T. Argyropoulos, J. Fuster, A. Vnuchenko, M. Boronat, M. Usó, O. García, C. Llácer
XBOX team: W. Wuensch, W. Wolley, G. McMonagle, N. Catalán-Lasheras, M. Volpi, I. Syratchev, A. Grudiev
Elettra guys: C. Serpico, S. Nuaman
To the Marie-Curie actions program

## Slide 33: THANK YOU FOR YOUR ATTENTION

## Slide 34: Data analysis: RF breakdown localization
More details in following talks.
We are interested in finding the regions which limit the overall performance of the structure. A combination of different methods is used.
Edge method: compares the time of detection in which the reflected power rises and the transmitted drops
Correlation method: compares the delay time between the incident signal and the repeated pattern of the reflected after breakdown forms.