## Slide 1
Institute for Instrumentation in Molecular Imaging, i3M
Valencia, October 27th 2025
Álvaro Anreus, Marta Freire, David Sánchez, Santiago Jiménez-Serrano, Edwing Y. Ulin-Barreño, Jorge Álamo, Julio Barberá, Neus Cucarella, Andrea González-Montoro, Alfonso Ríos, Irene Torres, Luis F. Vidal, Marc Gil, José M. Benlloch and Antonio J. González
IMAS: a Total-Body PET system with TOF and DOI capabilities
contact: aanrval@i3m.upv.es
V Jornadas RSEF / IFIMED de Física Médica

## Slide 2
Overview
Total Body PET: High Sensitivity Molecular Imaging (IMAS) project
Materials and Methods
System Performance
Conclusions and next steps

## Slide 3
Overview
Total Body PET: High Sensitivity Molecular Imaging (IMAS) project
Materials and Methods
System Performance
Conclusions and next steps

## Slide 4
PET systems
PET is a non-invasive, highly sensitive molecular imaging technique  functional information at the molecular level
But these studies are still limited by statistics, radiation dose and efficiency geometry
Total Body PET increases axial coverage and sensitivity.
Current PET scanners only collect about 1% of the emitted signals. TB-PET can increase sensitivity around x20.

## Slide 5
PennPET EXPLORER
TOF 250 ps
70 - 140 cm long
Philips technology platform
Pixels, no DOI
dSiPM Philips
$??
United Imaging uEXPLORER
Siemens Quadra
Highest sensitivity,
400 ps
2 m long
United imaging Healthcare
Pixels, no DOI
SiPM On-Semi
$$$$
TOF <214 ps
1.06 m long
Siemens Vision tech
Pixels, no DOI
SiPM Hama
$$$
J-PET
TOF impact
70 cm
Plastic scintillator
DOI
Total Body PET system
On the market
GE Omni
32 cm -> 96 cm
GE Omni
BGO
Pixels, no DOI
$$
IMAS
DOI
TOF 450 ps
71 cm long
Semi-monolithic LYSO
DOI
SiPM Hama
$$

## Slide 6
Total Body PET with DOI and TOF
Role of DOI
1.
DOI (Depth of Interaction) is the depth at which the annihilation photon interacts within the detector.
Role of TOF
TOF information  increases the image Signal to Noise Ratio (SNR)
With TOF: better localization of the true annihilation point along the LOR
TOF (Time of Flight): the difference in the arrival time of the two annihilation photons
With DOI: the true Line of Response (LOR) is determined
DOI information provide homogeneous
spatial resolution across the entire Field of View (FOV)
Non TOF
TOF (210 ps)

## Slide 7
Total Body PET: High Sensitivity Molecular Imaging (IMAS) project
Construction of a TB-PET system:
Long axial coverture (71 cm), 5 rings with a diameter of 82 cm and axial length of 10 cm each, separated by 5 cm between rings.
Without a high-cost impact: using semi-monolithic crystals and a signal reduction readout that significantly reduces the cost.
Detector technology with DOI and TOF capabilities: semi-monolithic crystal
IMAS, Imagen Molecular de Alta Sensibilidad

## Slide 8
Overview
Total Body PET: High Sensitivity Molecular Imaging (IMAS) project
Materials and Methods
System Performance
Conclusions and next steps

## Slide 9
Materials
Mini-Module (MM):
LYSO semi-monolithic coupled to 8x8 SiPM matrix
Reduction readout
PETsys electronics
FEM256
FEB/D 8k
System components:
16 x MM  1 Super-Module (SM), readout by 1 x FEM256
24 x SM  1 Ring
IMAS PET system  5 Rings: Axial length of 71 cm (5 cm gap)
  122 880 SiPMs
  30 720 channels
  15 360 slabs
Water cooling for temperature stabilization
Installation on May-June 2023
CT
PET
25.8 mm
25.8 mm
IMAS scanner

## Slide 10
IMAS Calibration
Spatial performance for the reference SM
 The trained NN are used to estimate the impact position in y-monolithic and DOI direction in all the SM of the IMAS system
Reference SM
Neural Network
Reference SM
Motor
SM for coincidences
Reference SM
22Na point sources
Experimental set-up
Slit collimator

## Slide 11
Methods
Raw data PETsys
LM
Reconstructed image
x- coordinates are estimated
y- and DOI coordinates are calculated using NN
Energy calibrated
DOI information included for correcting parallax errors
MLEM or OSEM
Normalization correction
Attenuation correction

## Slide 12
Overview
Total Body PET: High Sensitivity Molecular Imaging (IMAS) project
Materials and Methods
System Performance
Conclusions and next steps

## Slide 13
Spatial resolution
Point source 22Na, activity 6.2 µCi and 1 mm active diameter
MLEM reconstruction, 15 iter
cFOV
3/8 FOV
Point sources
No-DOI
DOI

## Slide 14
Sensitivity NEMA 2018
Pipe filled with FDG
Initial activity of 15.5 MBq
Axial position (mm)
Sensitivity (kcps/MBq)
Total Sensitivity:
56.5 kcps/MBq
Sensitivity phantom

## Slide 15
Count rate performance
Loss counts at:
~28 MBq
2023
2024
2025
Loss counts at:
~9.3 MBq
Loss counts at:
~44 MBq
Improvements 2024:
Installation of a second DAQ
SW/FW writing into disk performance increased: writing at 3 Gb/s into disk
FW implementation of a threshold on the minimum number of hits: remove noise events
Improvements 2025:
FW implementation of a threshold on the minimum energy: remove low energy events
 We are in discussions with PETsys to improve rate performance
Future improvements:
Improve write performance into disk:
limit is found at 6 Gb/s
Reduce coincidence window
2026

## Slide 16
Images: Torso NEMA 2018
Torso Phantom:
Ratio 1:10 FDG
OSEM reconstruction, 10 iter + 5 subset
Normalization applied
Attenuation corrected using CT

## Slide 17
Patients
Ethics committee approval September 2024. Room approval to use FDG February 2025.
Protocol is very “wide” allowing all recruited patients for conventional PET/CT to be asked for an IMAS PET/CT
Patient Protocol at Hospital La Fe
Patient injection
Patient consent obtained
TB-PET IMAS (no CT)
WB-PET/CT from La Fe

## Slide 18
Patients
RESULTS IMAS Patient 01 – 14/03/2025
Activity 3.59 mCi (FDG)
Time acquisition 800s

## Slide 19
Patients
RESULTS IMAS Patient 02 – 03/04/2025
Activity 3.41 mCi (FDG)
Time acquisition 1200s

## Slide 20
Patients: CT co-register
RESULTS IMAS Patient 02 – 03/04/2025
CT Philips -Radiant DICOM Viewer
CT Philips + PET IMAS
PET IMAS
CT Philips

## Slide 21
Overview
Total Body PET: High Sensitivity Molecular Imaging (IMAS) project
Materials and Methods
System Performance
Conclusions and next steps

## Slide 22
Conclusions
Spatial resolution along the center of FOV of scanner is around 3 mm with DOI correction, whereas without DOI correction it increases exponentially in the radial direction.
Sensitivity of the IMAS is around 56 kcps/MBq.
NECR performance is below expected, due to count rate limitations at the DAQ level and some limitations of the system design. The system currently loss counts at activity of 44 MBq.
Significant efforts are being made in collaboration with PETsys into data transmission bottle neck and rate capabilities of the system (a lot of events/s due to high sensitivity of the device).
First patient images acquired with IMAS system demonstrate excellent image quality and confirm the expected performance.
System | Axial FOV (cm) | Detector design | FWHM @20 cm (mm)
uExplorer1 | 200 | Pixel | 4.7 (BP)
PennPET2 | 140 (no gaps) | Pixel | 5.6 (OSEM)
Siemens Quadra3 | 106 | Pixel | 5.83 (BP)
IMAS | 70 (with gaps) | Semi-monolithic | 3.2 (MLEM)
1Spencer, B. A., et al., Journal of Nuclear Medicine, 62(6), 861-870,
2Dai, Bing, et al., Physics in Medicine & Biology 68.9 (2023): 095007,
3Prenosil, George A., et al., Journal of nuclear medicine 63.3 (2022): 476-484
Apply TOF in reconstructions
More patients
Nest steps

## Slide 23
Thank you for attention!
Institute for Instrumentation in Molecular Imaging, i3M
contact: aanrval@i3m.upv.es