## Slide 1: Model-Based Dose Calculation in Brachytherapy: Current Status and Future Directions
Gabriel Paiva Fonseca

## Slide 2: Summary
TG43
TG186
TPS commissioning
Reference database
Clinical results

## Slide 3: TG43
Simple formalism
Tabulated data
Easy to implement

## Slide 4: TG43
Robust … but a bit insensitive
Robust … can handle bad images

## Slide 5: TG186
This report addresses three major issues associated with the advent of MBDCAs, namely,
Choice of dose speciﬁcation medium (Sec. III), because there are energy-dependent differences even when non water materials are included in the radiation-transport model depending on whether dose is speciﬁed to a water reference medium or the local tissue medium.
Voxel-by-voxel cross-section assignment, as neither conventional single-energy CT nor ICRU/ICRP tissue composition recipes are de facto accurate guides for the task (Sec. IV).
Commissioning procedures for MBDCAs (Sec. V).

## Slide 6: Dose Speciﬁcation Medium

## Slide 7: Dose reporting

## Slide 8
Dose reporting
How should we convert Dm,m to Dw,m ?
It depends on the volume of interest

## Slide 9: Dose reporting
Even for 192Ir it can be relevant

## Slide 10: Voxel-by-voxel cross-section assignment

## Slide 11: Mat. Characterization and segmentation
Material characterization
CT/CBCT/DECT/Spectral segmentation
CT artefacts
Calcifications … Applicators … etc.
Contrast
Not common in brachytherapy … additional steps

## Slide 12: Commissioning TG186 and beyond

## Slide 13
The TG-186 report introduced two levels of commissioning:
Level 1 pertains to the veriﬁcation of the algorithm in reference conditions for the dosimetric characterization of a source according to TG-43. Tolerances < 2% (in dose or dose rate) are expected at clinically relevant distances.
Level 2 pertains to the comparison of the end-user’s MBDCA results with corresponding reference data in speciﬁc virtual phantoms mimicking clinical scenarios

## Slide 14: Workflow
Based on reference data
User most know limitations of the model

## Slide 15: What about now …
Moving to Zenodo soon

## Slide 16: GYN case
Coming:
  Head and Neck
  Prostate Iodine
In-preparation:
  Valencia/Leipzig
  Skin
  Neuro – Gamma tile
  60Co - Gyn
Case not intended to proof clinical relevance, but to provide a reference for a wide variety of scenarios

## Slide 17: Clinical use
Available TPS
  Oncentra ACE – Collapse cone
  BrachyVision – AcurosBV Deterministic approach
  RayStation - MC
Mostly in research setting
Very slow clinical adoption

## Slide 18: GYN
This study includes 32 patients with a total of 156 fractions treated with T&O applicators with titanium interstitial needles. ... The AcurosBV calculated doses are always lower than TG43 calculated dose. The average dose reduction was 4.20 % for target and 5.90 % for OARs. The maximal dose reduction was 9.68% for target and 29.03% for vaginal wall OAR.
Clinical implementation of a MBDCA AcurosBV style algorithm is challenging due to both logistical complexities and strain on computational resources. The results of this study do not present enough clinical dosimetric change to warrant using an MBDCA approach for the primary algorithm for HDR 192Ir T&O gynecological treatments. However, MBDCA approach could be utilized as a supplemental check or verification tool. Due to calculation speed, TG43 could be utilized as the primary algorithm with AcurosBV calculated dose as a second verification check.

## Slide 19: Lung
Lung – to be submitted
Large differences in low dose regions

## Slide 20: Liver
Up to ~7% V5Gy
Similar CTV coverage

## Slide 21: Where can it be more important ?
Low energy – Lacking TPS
Skin and shielded applicators
Needs to be added to the TPS library

## Slide 22: 3D printed applicators

## Slide 23: Where can it be more important ?

## Slide 24: Conclusion
Clinical data needed
TPS adoption is very slow ($$$)
Limited possibility in TPS for customization
New possibilities including for dosimetry (e.g.g energy correction)

## Slide 25: Thank you !